## Slide 1: MS2 as a Surrogate Assay

## Slide 2: Three experiments:
Effects of RT
Effects of untreated saliva
Standard Curve
Heat template at 95°C for 5 min
45°C for 5 min
95°C for 3 min
50 cycles
  95°C for 1 sec
  60°C for 10 sec

## Slide 3: RT
Template conc: 104 pfu/rxn
Primer conc: 0.9 uM
Probe conc: 0.25 uM
3 samples:
Add RT (positive control)
Do not add RT
Add RT but do not heat template

## Slide 4: RT Analysis
The +/- RT samples behaved as expected with the +RT showing a strong signal and
–RT not amplifying.
The sample without heating the template also shows a strong signal but has a higher cp suggesting that there is a mix of packaged and unpackaged RNA in the template.
+RT (23.14)
Packaged RNA (no heat) (25.16)
- RT

## Slide 5: Saliva
Template conc: 103 pfu/rxn
Primer conc: 0.9 uM
Probe conc: 0.125 uM
Add saliva to unpackaged RNA (heat template)
Add saliva to packaged RNA (do not heat template)
Positive Control (no saliva)
No RT
After combining saliva and template it was incubated at 37°C for 10 min then heat treated at 95°C for 10 min.

## Slide 6: Saliva Analysis
Positive Control (29.47)
Packaged RNA w/ saliva (33.13)
Unpackaged RNA w/ saliva
-RT
The saliva completely killed the rxn when added to unpackaged RNA but only partially killed it when added to packaged RNA.

## Slide 7: Standard Curve
105, 104, and 103, pfu/rxn
Primer concentration: 0.9 uM
Probe concentration: 0.25 uM

## Slide 8: Standard Curve Analysis
Efficiency = -1+10(-1/slope)
=1.809
1x103 (33.84)
1x105 (26.07)
1x104 (29.56)
NTC